## Slide 1: Brightness of Horizontal (H5) and Vertical (H1) Cold Sources
Ken Andersen

## Slide 2
H1
H5

## Slide 3: Extract Brightness from Neutron Measurement
Needs end-of-guide position with chopper and TOF analysis
  Readily available on H1: Figaro, D17, PF1B
  Not so readily available on H5 (IN15?)
Propagate backwards
Correct for attenuation through window materials etc.
Scale using Au-foil measurement

## Slide 4: Use spectrum measured on Figaro (H172)

## Slide 6: Detector spectrum corrected for chopper opening time
time
distance
Δt
Δt

## Slide 7: Al Bragg edges measured on D17

## Slide 8: Al window transmission
T=0.875
⇒t=20mm

## Slide 9: Al window transmission

## Slide 10: Detector efficiency
Window thickness = 10mm
- half of the total 20mm

## Slide 11: Air attenuation
Ageron measurement at 60% humidity
50cm through air

## Slide 12: Attenuation through frame-overlap mirror
0.7mm Si at 3° angle

## Slide 13: Final corrected spectrum
Flux (arb. units)

## Slide 14: Propagate backwards through guide

## Slide 15: Gold foil measurement
August 2007
after installation of H17 guides
at VS position: entrance into level C casemate
H171 (Figaro) Φcapt=1.07×1010 n/cm2/s
H172 (Granit) Φcapt=1.57×1010 n/cm2/s

## Slide 16: Evaluate capture flux at VS and rescale

## Slide 17: Comparison of capture fluxes
Capture Flux (n/cm2/s) | Measured | Calculated from Ageron | Present calculation
H171 (Figaro) VS | 4.24×1010 August 2007 | 2.88×1010 | (4.24×1010)
H172 (Granit) VS | 6.19×1010 August 2007 | 4.03×1010 | 5.98×1010
H172 (Granit) mono | 4.4-4.8×1010 April 2008 | 3.20×1010 | 4.63×1010

## Slide 18: Spectral flux at Granit monochromator
Spectral flux at λ=8.9Å (n/cm2/s/Å) | 
Previous estimate (2005-2007) | 6.4×108
Ageron brightness | 4.4×108
Present calculation | 5.4×108

## Slide 19: Brightnesses of HCS & VCS

## Slide 20: Summary
Previous estimate of 8.9Å flux at Granit was about right
Ageron VCS brightness is slightly low
  needs 50% readjustment in the 1.5-4Å range
  needs 20% readjustment around 9Å
HCS and VCS brightnesses are probably about the same (within 30%)